## Effective Thickness
Effective Thickness Calculation | [9 empty] | Cell to be manually filled with provided data to be entered in flashLine sw prior to the test
Cells already filled with theoretical value
Sample Holder ID |  |  |  |  | Effective Thickness |  | 0.11736 |  |  | Cells to be manually filled FlashLine tests results data
Measured Thickness |  | 0.0965 | [7 empty] | cells automatically filled with calculated data
Sample holder's effective thickness to be entered in FlashLine sw in future tests
Distilled Water Literature Value |  |  | 0.00145 |  | Ethylene Glycol Literature Value |  |  | 0.000939
Distilled Water Calibration Test Results |  |  |  |  | Ethylene Glycol Calibration Test Results
Test 1 |  |  |  |  | Test 1
Temperature | Shot 1 | Shot 2 | Shot 3 | Average | Temperature | Shot 1 | Shot 2 | Shot 3 | Average
20 | 0.00108 | 0.00101 | 0.00097 | 0.00102 | 20 | 0.00057834 | 0.0006227 | 0.00061975 | 0.00060693
20 | 0.00102 | 0.00116 | 0.0011 | 0.0010933 | 20 | 0.00062314 | 0.00062314 | 0.00062193 | 0.00062274
Test 2 |  |  |  |  | Test 2
Temperature | Shot 1 | Shot 2 | Shot 3 | Average | Temperature | Shot 1 | Shot 2 | Shot 3 | Average
20 | 0.00094791 | 0.0010308 | 0.0010129 | 0.00099718 | 20 | 0.0005988 | 0.00062634 | 0.00062484 | 0.00061666
20 | 0.0009854 | 0.0010251 | 0.000967 | 0.00099249 | 20 | 0.00060136 | 0.00058081 | 0.00059945 | 0.00059387
Test 3 |  |  |  |  | Test 3
Temperature | Shot 1 | Shot 2 | Shot 3 | Average | Temperature | Shot 1 | Shot 2 | Shot 3 | Average
20 | 0.00099 | 0.00113 | 0.00103 | 0.00105 | 20 | 0.00064016 | 0.00059385 | 0.00058467 | 0.00060623
20 | 0.0010294 | 0.00090573 | 0.0010273 | 0.00098746 | 20 | 0.00060438 | 0.00060421 | 0.00062115 | 0.00060991
Effective Thickness |  |  |  |  | Effective Thickness
| Test 1 | Test 2 | Test 3 | Average |  | Test 1 | Test 2 | Test 3 | Average
Segment 1 | 0.11506 | 0.11637 | 0.1134 | 0.11492 | Segment 1 | 0.12003 | 0.11908 | 0.1201 | 0.1198
Segment 2 | 0.11113 | 0.11664 | 0.11694 |  | Segment 2 | 0.1185 | 0.12134 | 0.11974

## Verification
Liquid Cell Verification Test
Sample holder's effective thickness to be entered in FlashLine sw prior to the test
Sample Holder ID |  |  |  |  | Effective Thickness |  | 0.11736 |  |  | Cells already filled with theoretical value
cells to be manually filled with tests results data
Distilled Water Literature Value |  |  | 0.00145 |  | Ethylene Glycol Literature Value |  |  | 0.000939 |  | cells automatically filled with calculated data
Distilled Water Calibration Test Results |  |  |  |  | Ethylene Glycol Calibration Test Results
Temperature | Shot 1 | Shot 2 | Shot 3 | Average | Temperature | Shot 1 | Shot 2 | Shot 3 | Average
20 | 0.00108 | 0.00101 | 0.00097 | 0.00102 | 20 | 0.00057834 | 0.0006227 | 0.00061975 | 0.00060693
20 | 0.00102 | 0.00116 | 0.0011 | 0.0010933 | 20 | 0.00062314 | 0.00062314 | 0.00062193 | 0.00062274
| Percent Error |  |  |  |  | Percent Error
| Shot 1 | Shot 2 | Shot 3 | Average |  | Shot 1 | Shot 2 | Shot 3 | Average
| 25.52% | 30.34% | 33.1% | 29.66% |  | 60.11% | 57.06% | 57.26% | 58.14%
| 29.66% | 20% | 24.14% | 24.6% |  | 57.02% | 57.02% | 57.11% | 57.05%